Treatment of  stable and unstable intertrochanteric fractures with selfdynamisable internal fixator / concept of double dynamisation
Introduction: Intertrochanteric fractures of femur are the third most common fractures among all bone fractures, and despite a number of different methods of treatment, no method has been applied not been able to present itself as the method of choice in the treatment of intertrochanteric fractures. Despite the improvement in the development of new implants, percentage of serious complications of the treatment of these fractures remains very high, varying in the range of 10 to 20%1-6. One of the most serious complications of internal fixation of intertrochanteric fractures is nonunion of fractures due to lack of additional axial dynamisation of implants. 
Material and methods: During the period from 2000 to 2009 we have analyzed the use of selfdynamisable internal fixator  (SIF implant) in the treatment of 247 patients with stable and unstable intertrochanteric fractures. Fracture types were classified according to the AO fracture classification/Orthopaedic Trauma Association scheme. Salvati-Wilson scoring systems being  used in our study for functional assessment consider pain, walking ability and hip movements operated patients. 
Results: Of the total number of treated patients, 134 were male and 113 females, aged 19 to 90 years (average 49.6). More than half of the patients were older than 50  years. Monitoring of patients after the operation was carried out clinically and radiographically for a period of three to six months in all patients, whereas in the two-year follow-up was conducted in a total of 176 patients (71,2%). Average duration of surgery was 47 minutes and the average blood loss of 145 ml, and the average fluoroscopy time was 16 s (8–97 s). The average time for union was 3.7 months (3 – 6,5 months). Double dynamisation (dynamisation along the neck and shaft of femur) was observed in 85 patients (34,4%), was on average 4.3 mm ( 1,5 - 8 mm). All fractures who underwent dynamisation implants have healed completely within no later than six months after surgery. In 17 cases there was a cut-out phenomenon of implant, while in seven cases was detected mechanical failure of implant. Complications were detected within 3 to 6 weeks after surgery, and treated by the method of intramedullary fixation. During the study, there were not detected any cases of infecton and  thromboembolic complications. 
Conclusion: The concept of double dynamisation improved the fracture healing in stable and unstable intertrochanteric fractures using the selfdynamisable internal fixator. This biological method of fixation provides healing of the intertrochanteric fracture in the optimum period of time, significantly reduce the risk for mechanical failure.
Key words: selfdynamisable internal fixator , stable and unstable intertrochanteric fractures, dynamisation
Introduction

The extension of life expectancy caused a growing number of patients with fractures of the proximal femur due to poor bone quality. Because of that osteosynthesis of intertrochanteric fractures is one of the most common operation in orthopedic trauma surgery. For the surgical treatment of these fractures still apply different implants, which can be roughly divided into extramedullary and intramedullary implants. Complication rates arising during treatment of these fractures using  different surgical techniques and principles ranges in the literature from 10 to 20%1-6. The most serious complication of treatment of trochanteric fractures are infection, the occurrence of hematoma, cut-out and cut-through. Also, complications in the form of delayed healing, nonunion and subsequent breakage of implants can have serious consequences for the patients7.
To avoid these complications, it is necessary to provide good fractures reduction, proper position of the lag screws, in order to achieve the most important biomechanical factors of stability of fixation. Very important biomechanical characteristics of each implant is the presence of dynamic components of the implant to provide a secondary impaction of the fracture to achieve union. If this biomechanical characteristics of the implant is not present,  in situations where there was delayed union of fracture may occur implants brakage as the implant becomes the weakest part of rigid fixation construct.
The use of dynamic implants with a sliding screw in the axis of the femoral neck is now the standard in the internal fixation of intertrochanteric fractures. Sliding screw allows controlled collapse (dinamysation) and physiological compression of bone fragments at the fracture site, while preserving the neck-shaft femur axis. Weight bearing on the injured leg leading to the telescopic movement head-neck part of the femur together with the sliding screw, so that the possibility of fracture dynamization directly depends on the ability of dynamization of the sliding screw. While standing on the injured leg, the head of the femur is exposed to the force resulting from the effects of body weight  and muscle contraction of the hip abductors8,9,10,11. This force acts downward and outward, and it tends to dynamization of  head-neck part of the femur with sliding screw and simultaneously performs cross slide screw load, leading to the appearance of sliding friction between the screw and the main part of the implant. This friction occurs at two points of support sliding screws of the implants: the lower part of the medial and upper lateral foramen through which the sliding screw passes. Dynamisation occurs only when overcome this friction12.
In our article we want to show validity of  application  of selfdynamisable internal fixator (SIF) in the treatment of stable and unstable intertrochanteric fractures.
Material and methods
Between January 2000 and December 2009 at the Department of Orthopedics and Traumatology in Nis, University hospital Nis, Serbia, in total 247 patients underwent treatment of intertrochanteric fracture of the femur using SIF metod of fixation. 134 were male and 113 females patients, aged 19 to 90 years (average 49.6). More than half of the patients were older than 50  years.  All fractures were classified according to AO/OTA system of fracture clasification. Those with pathological intertrochanteric fractures were excluded from the study. Our criterion  for distinguishing stable from unstable intertrochanteric fractures was  absence of medial support that was considered the main criterion for fracture instability13. According to this crierion,  97 consecutive patients in our study sustained an unstable intertrochanteric fracture of the hip (fracture types 31–A3.1-3,A2.2,A2.3) according to the AO classification.
All patients had their operation within 72 h of admission. The operation was carried out in operating room using an orthopaedic traction table and image intensifier. After closed reduction, the proximal femur was exposed through a lateral approach using 5–6-cm-long skin incision, beginning just distally from the trochanteric ridge, and a guide wire was passed into the femoral neck aiming at the centre of the femoral head using a 130 degree guide. The operation was carried out in a similar manner to a conventional  DHS technique using a biological reduction concept without attempt of open reduction or exposing the fracture . Radiographs were taken postoperatively and at 6 to 8-week intervals during follow-up until clinical and radiological healing was achieved. During the first 6 weeks postoperatively, patients were mobilised on partial to full weight bearing as far as were tolerated. Initial mobilisation was supervised by a physiotherapist. Mean hospital stay was 11.9 days (range 7– 25 days), depending on patient mobility and social circumstances.
During the follow-up visits, patients were assessed with respect to walking ability, hip movements and pain. Radiographs were taken to assess healing and neck-shaft angle.
[image: image1.emf]
Figure 1. The AO classification of trochanteric fractures.

Implant characteristics:

The main three characteristics of the selfdynamisable internal fixator (SIF) developed by Mitkovic, are: possibility of spontaneous axial dynamisation, preservation of both periosteal and medullary bone blood circulation, and less invasive technique of application. This implant was approved by the National Drug and Medical Devices Agency14,15. SIF has been investigated widely in laboratories Faculty of Mechanical Engineering and in terms of testing the maximum load to the destruction of the implant and the cyclic tests the effect of repetitive loading was also investigated experimentally in 60 animals during the preparation of a PhD thesis16.
The SIF (Traffix Ltd., Nis, Serbia) was made of stainless steel (ASTM F 138-2). There are three different lengths of  selfdynamisable internal fixator  for intertrochanteric fracture fixation, from 100 mm, 150mm, 200 mm, 250mm. This variant consists of a trochanteric unit (for dynamic hip screws), which extends distally as a bar. One or two clamps can be fixed to the bar. On the distal end, it has an anti-rotation dynamic unit, length 18mm. On the trochanteric unit, there are three holes, but it is enough to introduce two screws (7 mm diameter) only into the neck and head of the femur, at 130 degrees ( fig.2). 
[image: image2.jpg]


[image: image3.png]


[image: image4.emf]
Figure 2. Concept of double dynamisation with SIF implant
Results

Of the total number of treated patients, 134 were male and 113 females, aged 19 to 90 years (average 49.6). More than half of the patients were older than 50  years.  Average duration of surgery was 47 minutes and the average blood loss of 145 ml, and the average fluoroscopy time was 16 s (8–97 s). Of the total number of treated fractures (n=247), 150 patients with intertrochanteric fractures were classified as stable fractures (60,7%), while in 97 patients the fracture was classified as unstable (39,3%). Monitoring of patients after the operation was carried out clinically and radiographically for a period of three to six months in all patients, whereas in the two-year follow-up was conducted in a total of 176 patients (71,2%). The average time for union was 3.7 months (3 – 6,5 months). Double dynamisation (dynamisation along the neck and shaft of femur) was observed in 85 patients (34,4%), was on average 4.3 mm ( 1,5 - 10,5 mm). All fractures who underwent dynamisation implants have healed completely within no later than six months after surgery.
Cut-out phenomenon occurred in 17 cases (6,9%), while in seven cases was detected mechanical failure of implant(2,8%). These complications were detected within 3 to 6 weeks after surgery. They were ultimately treated by removing the implants and intramedullary fixation with long Gamma nail third generation without bone grafting. All revised intertrochanteric fractures healed safely. During the study, was not detected any cases of thromboembolic complications, not a single case of infection.
a)[image: image5.jpg]


 b) [image: image6.jpg]



c)[image: image7.jpg]




In cases of unstable intertrochanteric fractures, the average dynamisation was 6,3 mm ( 3,7 – 10,5 mm). Fracture consolidation had been achieved in the mean duration of 10,3 weeks ( 7-19 weeks). The patient started weight bearing as soon as possible after operation, with recommendation to weight bearing in the first three weeks as tolerated as well, and after that all patients started to walk full weight bearing with crutches.
here has beenwown re consolidationmisation of implant well, and after that all patients sartedher 1998;28:51-9.

Tech (Berl);199[image: image8.jpg]


[image: image9.jpg]


[image: image10.jpg]


[image: image11.jpg]



ččč

For functional assessment we used the Salvati and Wilson score, which showed excellent results in 175 fractures (70,1%), good results in 54 fractures (21.9%) and fair results in 18 (7.3%). Hardware removal was done in 157 patients, typically not before 12 months after fracture union.
Discussion
The role of dynamization is well known in the orthopedic profession as an essential component of healing fractures, especially of the lower extremities, as well as to promote fracture healing in an optimal timeframe. Fractures of the proximal femur, intertrochanteric fractures, require special attention during the treatment, not only because of its prevalence in the pathology of trauma, but also because of the high incidence of morbidity and mortality in injured patients. Therefore still ongoing debate in the professional literature on the selection of the appropriate implant for the treatment of intertrochanteric fractures of femur. The first implants used for the surgical treatment of intertrochanteric fractures were fixed-angle blade plates17,18. After using this device more than a decade, many authors recommended them as a treatment of choices for intertrochanteric femur fractures, but the main complaint was that these implants did not allow controlled collapse and impaction at the fracture site without penetration of the femoral head19. In the early seventies of the last century. dynamic hip screw (DHS) has begun to be used for the treatment of  intertrochanteric fractures more and more. One of the main reasons for the growing popularity of the application of this method of fixation was the ability of the implant to resist penetration and screw threads to increase fixation in the proximal fragment providing controlled spontaneous dynamisation of proximal fragment to achieve fracture healing. But, lack of axial dynamisation of this extramedullary implant has led to complications20 such as pulling off of the side plate from the femoral shaft and dissociation of the sliding compression hip screw from the barrel. Failure of DHS fixation is especially present during fixation of unstable intertrochanteric fractures, with failure rate range from 11% to 56%21,22. Despite this, sliding hip screw remains the method of choice in the surgical treatment of stable intertrochanteric femur fractures worldwide. In the largest published study to date of 1024 patients with intertrochanteric fractures treated with DHS fixation, the authors obtained low rate of fixation failure (3,2%), with the most common complications were cut-out of the lag screw from the femoral head with an incidence of 1.9%23. The analysis of our results after application of SIF method of fixation on 247 patients, retrospectively monitored, treatment complications were found in a total of 24 patients (9.72%). The major complicatiom was cut-out phenomen, wich has been detected in 17 cases (6,9%). The analysis of these cases showed that the incorrect lag-screws positioning into the femoral head led to this complication. Also, these cases with cut-out complications were detected in patients who underwent surgery at the beginning of the study, so the learning curve plays an important role in avoidance of technical errors of SIF method (fig. 5).
[image: image12.jpg]


 [image: image13.jpg]



Figure 5. Correct position of lag-screws SIF implant in AP and profile radiographic imaging to avoid cut-out complication

As an alternative to this extramedullary sliding device we devoleped SIF method of fixation for intertrochanteric femur fractures, at our institution (Clinic of orthopeadic and traumatology, University hospital, Niš, Serbia),  which has the possibility of dynamisation in two axes: axis of the femoral neck and the longitudinal femoral axis. After a period of biomechanical testing12 of implants and experimental use of laboratory animals16, the clinical application of this method we introduced in early 200014,15,24.  
In mid-1980 there is a development of the first intramedullary nails for fixation of intertrochanteric fractures, and the first clinical application of so-called Gamma Nail first generation  took  place in 1988, evolving from the concepts of Gerhard Küntscher to treat trochanteric fractures25. The intramedullary position of Gamma Nail and other new intramedullary implants (PFN,PFNA)  considered to provide better biomechanical stability and  still allowed control impaction of fracture26. Although considered to be mechanically stronger  to extramedullary implants, because they are closer to the mechanical axis of the femur and hence has a less bending moments on the implant, here is no consensus among surgeons which implant, extramedullary or intramedullary has a lower complication rate27. The increased risk of intraoperative and late fracture of femur and reoperation rate remain still problematic aspect of proximal femoral nailing28. The most recent randomised comparative study whose objective was to compare the clinical outcome and the rate of complications in the application of new design proximal  intramedullary implants, Gamma Nail and PFNA, showed that  the risk for experiencing a postoperative complication after Gamma3 nailing was 40% versus 45% after PFNA fixation, concluded that  there is no significant difference in the overall clinical outcome and risk of complications between the PFNA - and the Gamma 3 - treated patients during the first postoperative year. These complications were principally the impaction of the fracture area, the cut-out, and the fracture of the femoral shaft at the tip of the implant29. All proximal femoral intramedullary nails new generation  are designed to provide additional gliding mechanism of lag screw in the axis of the femoral neck, and dynamisation along the femoral shaft axis through oval holes in the distal part of the nail  for dynamic locking in both stable and unstable intertrochanteric fractures. In some cases, dynamisation of intramedullary proximal nail can be blocked so that the implant becomes the weakest point of fixation, leading to nail breakage. Complications in the form of delayed union, nonunion, and (subsequent) implant breakage are less frequent, but may also have devastating consequences for the patient. In this situation, R. Biber at al. recommended the operative technique so called ‘’lateral notching’’ - using a chisel they remove cortical bone right below the sleeve of the lag screw on the lateral side in order to achieve effective fracture impaction along the femoral shaft axis. Specifically, they noted in their series of 2369 patients with intertrochanteric fractures surgically treated with proximal femoral Targon nail, some cases (14 patients,0.6%) who were readmitted because of delayed union or nonunion. Normally distal dynamisation can be achieved either by removing the interlocking screws or by placing one screw into a long (‘dynamic’) interlocking hole of the nail, but in this particulary cases gliding of the nail was blocked by cortical support underneath the sleeve of the lag screw on the lateral side. In all detected cases healing was achieved 6 weeks later, after a new operation which included interlocking screw removal and lateral notching7. In our series operated patients with SIF method, dynamisation was observed in 85 patients (34,4%), where there was no need for a new subsequent surgery to achieve further impaction of the fracture fragments.
In conclusion, our clinical data recommended SIF as a safe extramedullary implants for the fixation of stable and unstable intertrochanteric fractures. It provides a stable biological fixation of proximal femoral fractures, further adding impaction of the fragments along each axis (the axis of the femoral neck and the axis of the femoral shaft) whenever it is necessary to achieve union.
Literatura SIF

1. Baumgaertner MR, Curtin SL, Lindskog DM Intramedullary versus extramedullary fixation for the treatment of intertrochanteric hip fractures. Clin Orthop Relat Res 1998;348:87–94

2. Boyd HB, Anderson LD  Management of unstable trochanteric fractures. Surg Gynecol Obstet 1961; 112:633–638

3. Fielding JW, Magliato HJ . Subtrochanteric fractures. Surg Gynecol Obstet. 1966;122:555–560

4. Simmermacher RK, Bosch AM, Van der Werken C. The AO/ASIF proximal femoral nail (PFN): a new device for the treatment of unstable proximal femoral fractures. Injury 1999; 30:327–332

5. Watson HK, Campbell RDJ, Wade PA. Classification, treatment and complications of the adult subtrochanteric fracture. J Trauma 1964; 4: 457–480

6. Werner-Tutschku W, Lajtai G, Schmiedhuber G, Lang T, Pirkl C, Orthner E Intra- and perioperative complications in the stabilization of per- and subtrochanteric femoral fractures by means of PFN. Unfallchirurg 2002; 105:881–885

7. Roland Biber ,Hermann-Josef Bail, Hans-Werner Stedtfeld. Lateral cortical notching in specific cases of delayed unions or nonunions after intertrochanteric and reversed fractures. Arch Orthop Trauma Surg 2013; 133: 495–501

8. Loch DA, Kyle,RF, Bechtold JE, et al: Forces required to initiate sliding in second-generation intramedulary nails. J Bone Joint Surg 1998;80:1626-31

9. Lengsfeld M, Stammberger U, Mokwa A, et al: Predicting load bearing of the hip joint. Computerized analysis with a 3-D multibody model of the human. Biomed Tech (Berl) 1994 Dec;39(12):307-12.

10. Krebs DE,Robbins CE, Lavine L, et al. Hip bio-mechanics during gait. J Orthop Sports Phys Ther 1998;28:51-9.

11. Ruszkowski I. Ortopedija.Zagreb,Hrvatska: Jugoslovenska medicinska naklada, 1979.

12. Milan M. Mitković,Miodrag T. Manić, Dušan Lj. Petković, Saša S. Milenković, Milorad B. Mitković : The force that  causes dynamization of the selfdynamisable internal fixator (SIF). Acta Chirurg Iugosl 2013;60(2); 87-91

13. Matthias Knobe MD, Gertraud Gradl MD, Andreas Ladenburger MD, Ivan S. Tarkin MD, Hans-Christoph Pape MD: Unstable Intertrochanteric Femur Fractures: Is There a Consensus on Definition and Treatment in Germany? Clin Orthop Relat Res (2013) 471:2831–2840

14. Milorad Mitkovic , S. Milenkovic , I. Micic D. Mladenovic , Milan Mitkovic : Results of the femur fractures treated with the new selfdynamisable internal fixator (SIF). Eur J Trauma Emerg Surg. April 2012;(2): 191-200

15. Mitković MB, Bumbasirević M, Milenković S, Micić ID, Mitković MM, Mitković MM, Mladenović DS, Todorović MD. Fractures of the upper part of the femur treated with Mitkovic selfdynamisable internal fixator (SIF). Acta Chir Iugosl. 2010;57:103–7.

16. Gajdobranski Dj. Influence of different methods of Internal Fixation on bone callus characteristics in experimental animals. Doctoral disertation, Medical Faculty,University of Nis (Serbia), 2004.

17. Holt EP. Hip fractures in the trochanteric region: treatment with a strong nail and early weight-bearing. J Bone Joint Surg 1963;45A:687-705.
18. Jewett EL. One-piece angle nail for trochanteric fractures. J Bone Joint Surg 1941;23:803-810.

19. Jacobs RR, Armstrong HJ, Whitaker JH, Pazell J. Treatment ofintertrochanteric hip fractures with a compression hip screwand a nail plate. J Trauma 1976;16:599-603.

20. Kulkarni SS, Moran CG. Results of dynamic condylar screw for subtrochanteric fractures. Injury. 2003;34:117–22.
 21. Haidukewych GJ, Israel TA, Berry DJ. Reverse obliquity fractures of the intertrochanteric region of the femur. J Bone Joint Surg Am May 2001;83-A(5):643–50. 
22. R. Willoughby . Dynamic hip screw in the management of reverse obliquity intertrochanteric neck of femur fractures. Injury, Int. J. Care Injured (2005) 36, 105-109.

23. N. Chirodian, Barbara Arch, Martyn J. Parker. Sliding hip screw fixation of trochanteric hipfractures: Outcome of 1024 procedures. Injury, Int. J. Care Injured (2005) 36, 793—800.

24. Ivan D. Micic, MD, Milorad B. Mitkovic, MD, Il-Hyung Park, MD, Desimir B. Mladenovic, MD, Predrag M. Stojiljkovic, MD, Zoran B. Golubovic, MD, In-Ho Jeon, MD. Treatment of Subtrochanteric Femoral Fractures  Using  Selfdynamisable Internal Fixator . Clinics in Orthopedic Surgery 2010;2:227-231.

25. Kuentscher G: Recent advances in the field of medullary nailing.Ann Chir Gynaecol Fenn1948,37:115
26. Curtis MJ, Jinnah RH, Wilson V, Cunningham BW: Proximal femoral fractures: a biomechanical study to compare intramedullary and extramedullary fixation.Injury 1994, 25(2):99-104.
27. Kjell Matre MD, Leif Ivar Havelin MD, PhD,Jan-Erik Gjertsen MD, PhD, Birgitte Espehaug MSc, PhD,Jonas Meling Fevang MD, PhD. Intramedullary Nails Result in More Reoperations Than Sliding Hip Screws in Two-part Intertrochanteric Fractures. Clin Orthop Relat Res (2013) 471:1379–1386.

28. Parker MJ, Handoll HH Gamma and other cephalocon​dylic intramedullary nails versus extramedullary implants for ext​racapsular hip fractures in adults. Cochrane Database Syst Rev CD000093, 2005.
29. J. Vaquero, J. Munoz, S.Prat, C.Ramirez, H.J.Aguado, E.Moreno, M.D. Perez. Proximal Femoral Nail Antirotation versus Gamma3 nail for intramedullary nailing of unstable trochanteric fractures. A randomised comparative study. Injury, Int. J. Care Injured 2012; 43 (S2), S47–S54.
Figure 3. 81-years old female patient 


a) with stable osteoporotic intertrochanteric fractures (AO type 31-A1.2); 


b) postoperative radiographs after osteosynthesis with selfdynamisable internal fixator; 


c) after 9 weeks  sufficient callus formation with double dynamisation of implant








Fig.4.a) Unstable, reverse obliquity, intertrochanteric fracture in 63 years old female patient (AO type 31-A3.3); b) postoperative radiograph after osteosynthesis  with selfdynamisable internal fixator; c) 6 weeks after operation, double dynamisation of implant has been achieved; d) After  twelve weeks radiographic image showing  a fracture healing

















